## Slide 1: A Modem Testbed for Body Area Network (BAN) Communication
Internship wrap-up presentation
      Yifang Chen
Mentor: Brian Taff
8-Aug-2017
Feb 9th, 2017
BIOTRONIK // CRM

## Slide 2: The present cardiac rhythm management (CRM) market
Traditional canned devices
    Generator implanted in chest
      Substantial scarring/bulge
    Transvenous lead connections to heart
      Infection risk
    Extensive testing/design for MRI safety
Leadless pacing
    Catheter-based implantation
      No chest scar
    Device implanted within heart
      Minimized infection risk
    Inherently MRI safe
The future
Notes: 60

## Slide 3: The value of BAN communication in CRM products
Multi-chamber leadless
    Beat-by-beat communication
      Few bits, sent often
    Enables synchronization of atrium and ventricle
      More physiologic heart function
Atrium
Ventricle
BAN
Home monitoring (HMSC)
Sub-Q hub
(BM)
    Leadless pacer stats relayed to web
      Many bits, sent daily
    Subcutaneous hub powers RF relay
      Returns core canned
      device functionality
Anti-tachycardia pacing (ATP)
    Extends capabilities of new-generation defibrillators
      Few bits, sent rarely
    Less painful, low power means for halting some fatal rhythms
BAN
RF
GSM
Shock box
Leadless pacer
BAN
(Please pace)
Notes: 60

## Slide 4: Project Overview
Outline:
Develop library of SW-defined radio modems
Build infrastructure for physical-layer testing
Collect test result
Document findings
Software-defined Radio Bench Modems
Transmitter
Receiver
Galvanic
Acoustic
Data
Data
OR
Digital
Analog
Acoustic
Galvanic
Physical-layer test environment
Notes: 90s

## Slide 5: Project Overview
Outline:
Develop library of SW-defined radio modems
Build infrastructure for physical-layer testing
Collect test result
Document findings
Software-defined Radio Bench Modems
Physical Testing Environment
Transmitter
Receiver
Galvanic
Galvanic
Acoustic
Acoustic
Or
Analog
Digital
Data
Data
Notes: Add icon

## Slide 6: Develop library of SW-defined radio modems
Packet Encoder/Decoder
    A translator between human language and computer language
    text   binary digits
Interface to Transmitter/Receiver
    Tell the transmitter/receiver how I want to send/receive the modulated signal
Signal Modulation/Demodulation
    A translator between computer language and physical world language
    binary digits   physical wave
The targeted text message
Access code
‘hello
world’
Code to indicate the beginning of packet
10101100…001 (varies)
01001000 01100101 01101100 01101100 01101111 01110111 01101111 01110010 01101100 01100100
ASK:
Amplitude shift
FSK:
Frequency shift
PSK:
Phase
shift
…
Tx
Tx
Rx
Rx
Notes: 90s:
1\why we need to develop a system?
2\workflow

## Slide 7
ASK
FSK
PSK
Frequency Shift Key:
Good noise immunity
More bandwidth efficient
More complex design than ASK
0
1
Phase Shift Key:
Good noise immunity
Slightly more bandwidth efficient
More complex design than FSK
More Power Consumption
Amplitude Shift Key:
Bad noise immunity
Less bandwidth efficient
Simple design
1
0
0
1
Time Domain
Constellation Domain
Frequency Domain
Frequency Domain
Frequency Domain
------------------------- binary digits 1,0
------------------------- modulated signal
25KHz
35KHz
Time Domain
Notes: 60s

## Slide 8: A Deeper Look into FSK
Modulation (TX)
De-modulation (RX)
Notes: 60s

## Slide 9: A Deeper Look into FSK
Modulation(TX)
De-modulation (RX)
Notes: Picture need to change; need VCO?

## Slide 10: A Deeper Look into FSK
Modulation
De-modulation (RX)
Message interval: only noise
Message interval: only noise
Receiver NOT prepared for sudden jump from noise to message
Fill empty interval with 101010… preamble sequence
Access code + “helloworld”
Receiver prepared for receiving targeted message by 1010…
Message interval: only noise
Message interval: only noise
Notes: 30s

## Slide 11: A Deeper Look into FSK
Modulation(TX)
De-modulation (RX)
Notes: Picture need to change; need VCO?

## Slide 12: A Deeper Look into FSK
‘0’  25kHz
‘1’  35KHz
Modulation(TX)
De-modulation (RX)
Notes: Picture need to change; need VCO?

## Slide 13: A Deeper Look into FSK
Up-sampling:
Repeat each symbol X 200 times
‘0’  25kHz
‘1’  35KHz
Modulation(TX)
De-modulation (RX)
Notes: 30s

## Slide 14: A Deeper Look into FSK
Up-sampling:
Repeat each symbol X 200 times
‘0’  25kHz
‘1’  35KHz
Modulation
De-modulation (RX)
Voltage Controlled Oscilloscope:
(1) Generate Cosine Waves; (2) Control its frequency using input voltage
Signal with 25K (0) and 35K (1) volts
25000
35000
…
Amplitude
1
0.5
0
-0.5
-1
Amplitude
Notes: 30s

## Slide 15: A Deeper Look into FSK
Up-sampling:
Repeat each symbol X 200 times
‘0’  25kHz
‘1’  35KHz
Modulation(TX)
De-modulation (RX)
Notes: 30s

## Slide 16: A Deeper Look into FSK
Up-sampling:
Repeat each symbol X 200 times
‘0’  25kHz
‘1’  35KHz
Modulation(TX)
De-modulation (RX)
Notes: Picture need to change; need VCO?

## Slide 17: A Deeper Look into FSK
Up-sampling:
Repeat each symbol X 200 times
‘0’  25kHz
‘1’  35KHz
Modulation
De-modulation
0
1
DC offset (don’t matter)
Filtered out the high frequency noise
(for better understand, here I use frequency domain instead od time domain before)
noise
noise
Down-sampling: /200
Received Signal in Frequency Domain
Low passed signal in Frequency Domain
Notes: 30s

## Slide 18: A Deeper Look into FSK
Up-sampling:
Repeat each symbol X 200 times
‘0’  25kHz
‘1’  35KHz
Down-sampling:
/100
Modulation(TX)
De-modulation (RX)

## Slide 19: A Deeper Look into FSK
Up-sampling:
Repeat each symbol X 200 times
‘0’  25kHz
‘1’  35KHz
Down-sampling:
/100
Modulation
De-modulation
35KHz ->positive value (1); 25KHz ->negative value (0);
0
1
DC offset (don’t matter)
Low passed signal in Frequency Domain
Recovered Signal in Time Domain
Example here is 1,0,1,0,… signal sequence
Notes: 30s

## Slide 20: A Deeper Look into FSK
Up-sampling:
Repeat each symbol X 200 times
‘0’  25kHz
‘1’  35KHz
Modulation(TX)
De-modulation (RX)
Down-sampling:
/100
Notes: Picture need to change; need VCO?

## Slide 21: A Deeper Look into FSK
Up-sampling:
Repeat each symbol X 200 times
‘0’  25kHz
‘1’  35KHz
Up-sampling:
/200
‘0’  25kHz*2pi; ‘1’  35KHz*2pi
Modulation
De-modulation
    21
(1)Down-sample again:/2 to recover the original symbol rate
(2) Further Smooth the signal
Example here is 1,0,1,0,… signal sequence
Recovered Signal in Time Domain
Further Recovered Signal in Time Domain
Notes: Picture need to change; need VCO?

## Slide 22: A Deeper Look into FSK
Up-sampling:
Repeat each symbol X 200 times
‘0’  25kHz
‘1’  35KHz
Modulation(TX)
De-modulation (RX)
Up-sampling:
/200

## Slide 23: A Deeper Look into FSK
Up-sampling:
Repeat each symbol X 200 times
‘0’  25kHz
‘1’  35KHz
Up-sampling:
/200
/ 2
Modulation
De-modulation
    23
Truncate float to Binary:
positive 1; negative 0
Example here is 1,0,1,0,… signal sequence
1
0
Notes: 20-23: 30s

## Slide 24: Project Overview
Outline:
Develop library of SW-defined radio modems
Build infrastructure for physical-layer testing
Collect test result
Document findings
Software-defined Radio Bench Modems
Physical-layer Testing
Transmitter
Receiver
Or
Digital
Analog
Galvanic
Galvanic
Acoustic
Acoustic
Data
Data
Notes: inkspace digital->analog; separate two iLP;

## Slide 25: Project Overview
Outline:
Develop library of SW-defined radio modems
Build infrastructure for physical-layer testing
Collect test result
Document findings
Software-defined Radio Bench Modems
Physical-layer Testing
Transmitter
Receiver
Data
Data
Or
Digital
Analog
Galvanic
Galvanic
Acoustic
Acoustic
Notes: inkspace digital->analog; separate two iLP;

## Slide 26: Physical Testing Environment
  Distance: 40mm – 200 mm
  Orientation:
Right Atrium (RA)
Right Ventricle (RV)
BM chest
S-ICD
~70mm
Right Ventricle (RV)
~140mm
Right Atrium (RA)
Or
Right Ventricle (RV)
Right Atrium (RA)
Or
~130mm
Skin Simulation: Anechoic chamber (container surrounded by anechoic foam)
Body Fluid Simulation:
Saline concentration = 2.0g/L; Conductivity 2.85mS/cm
Air in lungs Simulation:
Using Balloon
only needed in S-ICD case
salt
BAN Testing Parameters:
  Transmission Media:
Notes: 90s

## Slide 27: Physical Testing Environment
Device positioning fixture:
Enables stable device positioning
Actual use in saline tank

## Slide 28: Project Overview
Outline:
Develop library of SW-defined radio modems
Build infrastructure for physical-layer testing
Collect test result
Document findings
Software-defined Radio Bench Modems
Physical-layer Testing
Transmitter
Receiver
Or
Digital
Analog
Galvanic
Galvanic
Acoustic
Acoustic
Data
Data
Notes: inkspace digital->analog; separate two iLP;

## Slide 29: Project Overview
Outline:
Develop library of SW-defined radio modems
Build infrastructure for physical-layer testing
Collect test result
Document findings
Software-defined Radio Bench Modems
Physical-layer Testing
Transmitter
Receiver
Or
Digital
Analog
Galvanic
Galvanic
Acoustic
Acoustic
Data
Data
Notes: inkspace digital->analog; separate two iLP;

## Slide 30: Outline for testing dual-chamber use case
Rx/Tx iLP
Orientation
Distance
Acoustic
Galvanic
Parallel
Cross
Face-to-face
100mm
Modulation Methods
Total Required Bandwidth
20KHz
FSK
Noise
No noise
Noise (f = carrier + 9KHz)
HW Factors
SW/Algorithm-factors
Environmental Factors
Accuracy:
1 Byte Message
10 Byte Message
Parameter Scope
Testing Values
Parameters
1.Received packet percentage
2.Bit Error Rate (BER) of received packets
3.Estimated total Bit Error Rate
Notes: 60s

## Slide 31: Testing Result – Screenshot example
Original Signal VS Demodulated Signal
“Hello World!”
“H@llo Werld!”
RX
TX
Bit Error Rate ~= 35%
Notes: 30s

## Slide 32: Preliminary test results using a galvanic physical layer
How does noise influence accuracy?
BER rises exponentially with increasing noise – for a fixed distance, location, and packet size
~200mV = max. permissible noise voltage for 1990KHz signal (signal amplitude ~= 1.54V)
Noise voltage (mV)
Bit Error Rate
Notes: 90s:
1\why we need to develop a system?
2\workflow

## Slide 33: Testing Result – Parameters that influence galvanic accuracy
How orientation influence accuracy? (each series corresponds to a unique distance)
For face-to-face orientation, larger distance will give more accurate transmission at least in range around 100mm
-----: smaller distance (use middle as reference point)
-----: larger distance (use head as reference point)
Bit Error Rate
Noise voltage (mV)
Face-to-Face Orientation
Noise voltage (mV)
Cross Orientation
Notes: 90s:
1\why we need to develop a system?
2\workflow

## Slide 34: Preliminary test results using a galvanic physical layer
How does orientation influence accuracy?(each series corresponds to each orientation)
General trend for improved data transception: Parallel > Face-to-face > Cross
-----: Face-to-Face
-----: Parallel
Bit Error Rate
Noise voltage (mV)
-----: Cross
Notes: 90s:
1\why we need to develop a system?
2\workflow

## Slide 35: Summary
Provided SDD with a new configurable tool for assessing BAN modem design templates and performing comparative testing of physical layer communication technologies
  Includes code (SW in GNU radio) that is readily portable to more mature test hardware (i.e. National Instruments vs. HackOne RF)
  Interfaces with a saline test environment capable of mimicking a multitude of BAN communication use cases
Begin preliminary testing that should ultimately facilitate the first “apples to apples” comparisons of in-house SDD physical layer communication methodologies
Documented my development and design efforts to enable support for continued system-level benchmarking after my tenure
Notes: 30s

## Slide 36: Possible project extensions
  Further improve the assessment system
    Develop additional modulation templates inclusive of PPM
    Transform the high-level SDR system into FPGA language to better assess Tx/Rx efficiencies
    Transition targeted design blocks to analog implementations to better mimic envisionable implant module embodiments
Collecting more testing data
    Explore full suite of developed modem formats and their impact on Tx/Rx accuracy
    Obtain better insight on the role that distance and orientation play
    Extract power efficiency metrics to shed light on design trade-offs